## Control
Field | Tomas baseline | QC final
Prepared by | Tomas Li, QC Stability | Viktor Takacs, QC technical reviewer
Prepared | 2027-02-06 | 2027-02-08
Study | STB-AVI401-002 | STB-AVI401-002
Decision package | 24-month pull, testing, and authority response | 24-month pull, testing, and authority response
QC decision owner | Viktor Takacs | Viktor Takacs
Required decision date | 2027-02-08 | 2027-02-08
Tomas recommendation | Authorize the 1 March pull and the full campaign. Use the current calibration package and routine due-month conventions. Lock the primary dataset for the 20 March response and append late CCI data to COM-AVI401-008. | QC does not adopt the baseline recommendation. Authorize the scheduled pull into controlled quarantine; do not begin reportable testing until the CHM-25-03 qualification interval and sample impact are resolved. Thereafter, only tests with a demonstrated qualified chain may proceed; assay, impurities, osmolality and CCI have additional holds.
QC disposition |  | CONDITIONAL APPROVAL — scheduled pull into controlled quarantine may proceed. All testing is held at the storage-chain gate; after that gate closes, testing remains split GO/HOLD by test. No complete-dataset or commitment handoff until every required 24-month result is verified.
Restrictions that remain active |  | HOLD: all reportable testing pending CHM-25-03 qualification/impact resolution; additionally osmolality (RS-OSM-004 certificate absent), assay (RS-AVR-014 absent), impurities (RS-AVR-014 and MTH-HPLC-012 revalidation absent), CCI (CCI-02 calibration absent); dataset release; unsupported response claims; MAT-OWR-005.
Next attributable QC review trigger |  | First trigger: Validation-approved CHM-25-03 qualification-status and sample-impact evidence (ACT-12). Then QC reviews each test-specific item (ACT-02 to ACT-05), the 2027-03-19 response pack, and the complete verified 24-month dataset.
Note on C9: QC-01 — This is an operating authorization for the working package only. It does not update LIMS, RIM, DMS, QMS or any other governed system.

## Operation Gates
Operation | Planned date | Tomas baseline decision | Baseline dependency | QC decision | QC basis / correction | Accountable owner | Supporting function | Due / gate | Objective closure evidence | Controlled source locator | QC comment ref
Physical pull and receipt | 2027-03-01 | GO | CHM-25-03 calibration current through 2027-12-05 | GO / QUARANTINE | Proceed with the fixed 2027-03-01 pull to protect the scheduled timepoint, but place samples in controlled quarantine. Current calibration covers the date; the reviewed chamber requalification evidence does not cover 2027 (36-month interval elapsed). | QC Stability Scientist | Stability Operations | 2027-03-01 08:00 | Contemporaneous pull/receipt record for all lots plus controlled quarantine record; no reportable testing before ACT-12 closes. | Study Status, Pull Schedule rows 20:22; AVT-STB-102-A4 v2.0 section 3; CC-2023-004; MER-2026-4587; ACT-12 | QC-02
Appearance / pH | 2027-03-02 to 2027-03-04 | GO | Routine bench sequence | CONDITIONAL GO | Method/test assets are evidenced current, but reportable testing may start only after Validation/QC closes the CHM-25-03 storage-chain impact gate (ACT-12). | QC Stability Scientist | QC Analytical | After ACT-12; then before first test | ACT-12 closed, then pre-use chain check and valid independently reviewed results for all lots. | AVT-SPEC-018 v03; MTH-VIS-002 v3.0; MTH-PHY-004 v3.0; AVT-VAL-358; AVT-VAL-348; MER-2026-7741; AVT-RS-101 | QC-03
Osmolality / particulate matter | 2027-03-08 to 2027-03-12 | GO | Current instruments and standards | PARTIAL CONDITIONAL GO | After ACT-12 closes, particulate matter may proceed. Osmolality remains held until both ACT-12 and RS-OSM-004 evidence (ACT-02) close. | QC Stability Scientist | QC Reference Standards Coordinator | After ACT-12; osmolality also after ACT-02 | ACT-12 closure; particulate chain check/results; osmolality also needs effective RS-OSM-004 evidence. | AVT-SPEC-018 v03; MTH-PHY-015 v3.0; MER-2026-9014; AVT-RS-084; MTH-PHY-009 v3.0; ACT-02 | QC-04
Assay | 2027-03-15 to 2027-03-18 | GO | RS-AVR-014 effective 2025-05-03 | HOLD | HOLD for missing RS-AVR-014 and for the common storage-chain gate. HPLC-04 and MTH-HPLC-011 are otherwise current. | QC Reference Standards Coordinator | QC Stability Scientist | Evidence accepted before planned start 2027-03-15 | ACT-12 closed plus effective RS-AVR-014 certificate covering use dates and documented pre-use chain check. | MTH-HPLC-011 v3.0; AVT-VAL-350; MER-2026-8014; ACT-03 | QC-05
Impurities | 2027-03-22 to 2027-03-26 | GO | MTH-HPLC-012 current through June 2027 | HOLD | HOLD for missing MTH-HPLC-012 revalidation and RS-AVR-014, and for the common storage-chain gate. HPLC-04 and impurity-standard certificates are otherwise current. | QC Method Owner | QC Reference Standards Coordinator; QC Stability | Evidence accepted before planned start 2027-03-22 | ACT-12, ACT-03 and ACT-04 closed, followed by documented pre-use chain check. | MTH-HPLC-012 v3.0; MER-2026-8014; AVT-RS-079; AVT-RS-076; ACT-03/ACT-04 | QC-06
Sterility | 2027-03-02 to 2027-03-16 | GO | Routine incubation | CONDITIONAL GO | Method and incubator are evidenced current, but reportable testing may start only after ACT-12 closes. | QC Microbiology Manager | QC Stability Scientist | After ACT-12; then before first test | ACT-12 closed, then pre-use chain check, valid controls, completed incubation and independently reviewed results. | AVT-SPEC-018 v03; MTH-MIC-001 v3.0; AVT-VAL-347; MER-2026-9221 | QC-07
Container closure integrity | 2027-04-13 to 2027-04-14 | GO | CCI-02 due April 2027 | HOLD | HOLD for missing CCI-02 calibration evidence and for the common storage-chain gate. Method revalidation is current. | QC Instrument Owner | QC Stability Scientist | Evidence accepted before planned start 2027-04-13 | ACT-12 and ACT-05 closed, then valid run and independently reviewed results for all lots. | AVT-SPEC-018 v03; MTH-CCI-003 v3.0; AVT-VAL-338; ACT-05 | QC-08
Complete dataset lock / QC release | 2027-03-19 | GO | Primary results complete; late CCI may follow | HOLD | The proposed 2027-03-19 lock predates planned impurity testing and CCI. A complete 24-month dataset requires all specification tests, including sterility and CCI, with raw-data and calculation review. | QC Stability Scientist | QC Analytical; QC Microbiology | After last required result; target review 2027-04-16 | All 30 required lot/test results reportable and independently reviewed; deviations resolved; completeness reconciliation approved. | AVT-SPEC-018 v03 section 2; AVT-SOP-024 sections 7-8; ACT-14 | QC-09
Written-response evidence handoff | 2027-03-19 | GO | Use complete 24-month package | CONDITIONAL GO | Handoff by 2027-03-19 may include only verified evidence already available and the independently reviewed 18-month regression. It must state that 24-month results do not yet exist and must not represent the package as the complete 24-month dataset. | Regulatory CMC Manager | QC Stability; MSAT; Validation | 2027-03-19 (response due 2027-03-20) | Point-by-point evidence matrix with precise citations, explicit unavailable-data statements, resolved ACT-15 to ACT-19, QC check and Regulatory approval. | HA-AVI401-IR-2027-014 pp.1-2; AVT-SOP-072; ACT-15:ACT-19 | QC-10
COM-AVI401-008 verified-data handoff | 2027-04-29 | GO | Reuse the response package | HOLD | COM-AVI401-008 is the next verified 24-month stability update, not a reuse of the March response. Handoff requires the complete verified dataset, including CCI, plus corrected trend interpretation. | Regulatory CMC Manager | QC Stability Scientist | After QC dataset release; target 2027-04-29 | QC-approved complete 24-month data package and trend assessment linked to COM-AVI401-008; Regulatory acceptance recorded in RIM by its owner. | HA-AVI401-IR-2027-014 point 7; Study Status Study B8; ACT-20 | QC-11
Note on E2: QC-02 — Pull is authorized as a sample-protection step, not as evidence that the storage chain was qualified. Samples remain quarantined pending ACT-12.
Note on E3: QC-03 — Test-specific chain is acceptable, but the chamber is part of every result chain under AVT-STB-102-A4; ACT-12 is a common prerequisite.
Note on E4: QC-04 — The combined row is intentionally split after the common storage gate: particulate matter may proceed; osmolality has an additional standard hold.
Note on E7: QC-07 — Sterility has an acceptable test-specific chain but remains blocked by the common storage qualification/impact gate.
Note on E9: QC-09 — A 19 March complete-data lock is impossible because impurity testing is planned for 22–26 March and CCI for 13–14 April.
Note on E10: QC-10 — The authority explicitly permits use of existing records and says unavailable measurements must be identified as unavailable.
Note on E11: QC-11 — The authority asks for verified 24-month results as the next commitment update, not as an attachment that predates testing.

## Qualified Chain
Specification test | Method | Instrument | Reference standard | Tomas baseline state | QC decision | QC basis / correction | Accountable owner | Supporting function | Due / gate | Objective closure evidence | Controlled source locator | QC comment ref
Appearance | MTH-VIS-002 | UV-02 backup | N/A | Current | GO | Correction: MTH-VIS-002 states no instrument and no reference standard; UV-02 is outside this chain. Method revalidation 2025-06-12 is within the 60-month interval. | QC Stability Scientist | QC Analytical | Before use 2027-03-02 | Effective method and documented no-instrument/no-standard pre-use check. | MTH-VIS-002 v3.0 sections 2-3; AVT-VAL-358 | QC-12
pH | MTH-PHY-004 | PH-03 | Not listed | Current | GO | Correction: method-required RS-BUF-021 was omitted from baseline. PH-03 calibration 2026-09-20 and RS-BUF-021 expiry 2027-10-11 cover planned use; method revalidation 2025-02-10 is current. | QC Stability Scientist | QC Reference Standards Coordinator | Before use 2027-03-02 | Pre-use verification of MTH-PHY-004, PH-03 and RS-BUF-021. | MTH-PHY-004 v3.0; AVT-VAL-348; MER-2026-7741; AVT-RS-101 | QC-13
Assay | MTH-HPLC-011 | HPLC-04 | RS-AVR-014 | Current to 2027-05-03 | HOLD | MTH-HPLC-011 and HPLC-04 are current, but no controlled certificate/expiry for required RS-AVR-014 was located. Baseline effective date is not an expiry determination. | QC Reference Standards Coordinator | QC Stability Scientist | Before assay start 2027-03-15 | Effective RS-AVR-014 certificate covering use date. | MTH-HPLC-011 v3.0; AVT-VAL-350; MER-2026-8014; ACT-03 | QC-14
Impurity A | MTH-HPLC-012 | HPLC-04 | RS-AVR-014 | Current to 2027-05-03 | HOLD | Correction: MTH-HPLC-012 requires RS-AVR-014, RS-IMPA-009 and RS-IMPB-006, not only RS-AVR-014. HPLC-04 and impurity-standard certificates are current; method revalidation and RS-AVR-014 evidence are absent. | QC Method Owner | QC Reference Standards Coordinator | Before impurity start 2027-03-22 | Effective MTH-HPLC-012 revalidation and RS-AVR-014 certificate; pre-use chain check includes all three standards. | MTH-HPLC-012 v3.0; MER-2026-8014; AVT-RS-079; AVT-RS-076; ACT-03/ACT-04 | QC-15
Impurity B | MTH-HPLC-012 | HPLC-04 | Not listed | Current | HOLD | Correction: same complete chain as Impurity A. RS-IMPB-006 is current to 2028-01-31; MTH-HPLC-012 revalidation and RS-AVR-014 evidence remain absent. | QC Method Owner | QC Reference Standards Coordinator | Before impurity start 2027-03-22 | Same objective closure evidence as row 5. | MTH-HPLC-012 v3.0; AVT-RS-076; ACT-03/ACT-04 | QC-16
Total impurities | MTH-HPLC-012 | HPLC-04 | Not listed | Current | HOLD | Correction: total impurities uses the same MTH-HPLC-012 chain and all three method-cited standards. Missing method revalidation and RS-AVR-014 evidence prevent use. | QC Method Owner | QC Reference Standards Coordinator | Before impurity start 2027-03-22 | Same objective closure evidence as row 5. | MTH-HPLC-012 v3.0; ACT-03/ACT-04 | QC-17
Osmolality | MTH-PHY-009 | OSM-02 | RS-OSM-004 | Current | HOLD | OSM-02 calibration and MTH-PHY-009 revalidation are current; no controlled certificate/expiry for required RS-OSM-004 was located. | QC Reference Standards Coordinator | QC Stability Scientist | Before use 2027-03-08 | Effective RS-OSM-004 certificate covering use date. | MTH-PHY-009 v3.0; AVT-VAL-344; MER-2026-8830; ACT-02 | QC-18
Particulate matter | MTH-PHY-015 | PM-01 | Not listed | Current | GO | Correction: MTH-PHY-015 requires RS-PSL-012. PM-01 calibration 2026-12-01, standard expiry 2028-06-02 and method revalidation 2024-08-22 cover planned use. | QC Stability Scientist | QC Analytical | Before use 2027-03-08 | Pre-use verification of method, PM-01 and RS-PSL-012. | MTH-PHY-015 v3.0; AVT-VAL-355; MER-2026-9014; AVT-RS-084 | QC-19
Container closure integrity | MTH-CCI-003 | CCI-02 | N/A | Current in due month | HOLD | Method revalidation is current, but no controlled CCI-02 calibration evidence was located. “Current in due month” does not demonstrate qualification on the use date. | QC Instrument Owner | QC Stability Scientist | Before use 2027-04-13 | Effective CCI-02 calibration certificate covering 2027-04-13 to 2027-04-14. | MTH-CCI-003 v3.0; AVT-VAL-338; ACT-05 | QC-20
Sterility | MTH-MIC-001 | INC-03 | N/A | Current | GO | MTH-MIC-001 revalidation 2024-01-18 and INC-03 calibration 2026-12-19 are within the appendix intervals; method cites no reference standard. | QC Microbiology Manager | QC Stability Scientist | Before use 2027-03-02 | Pre-use verification of method and INC-03; valid growth controls. | MTH-MIC-001 v3.0; AVT-VAL-347; MER-2026-9221 | QC-21
Storage | Protocol appendix | CHM-25-03 | N/A | Calibration demonstrates qualification | HOLD | Correction: calibration alone does not establish qualification. The reviewed requalification change completed in 2023; under the appendix 36-month interval it does not cover March 2027. Current calibration does not cure the qualification gap. | Validation Engineer | Stability Operations; QA | Before any reportable 24-month testing | Approved current requalification evidence covering storage, or QA-approved impact assessment/disposition for samples stored after interval expiry, with any required deviation linkage. | AVT-STB-102-A4 v2.0 section 3; CC-2023-004; AVT-VAL-316; MER-2026-4587; ACT-12 | QC-22
Contingency assets | Various | GC-01 / KF-01 | N/A | Include in clearance | N/A | Correction: GC-01 and KF-01 serve methods outside AVI401-SPEC-v03 and are excluded by the appendix definition; they are not campaign-clearance prerequisites. | QC Stability Scientist | None | No gate | Specification-to-method reconciliation showing no cited method uses GC-01 or KF-01. | AVT-STB-102-A4 v2.0 section 2; AVT-SPEC-018 v03 section 2 | QC-23
Note on F2: QC-12 — Correction: MTH-VIS-002 states no instrument and no reference standard; UV-02 is outside this chain. Method revalidation 2025-06-12 is within the 60-month interval.
Note on F3: QC-13 — Correction: method-required RS-BUF-021 was omitted from baseline. PH-03 calibration 2026-09-20 and RS-BUF-021 expiry 2027-10-11 cover planned use; method revalidation 2025-02-10 is current.
Note on F4: QC-14 — MTH-HPLC-011 and HPLC-04 are current, but no controlled certificate/expiry for required RS-AVR-014 was located. Baseline effective date is not an expiry determination.
Note on F5: QC-15 — Correction: MTH-HPLC-012 requires RS-AVR-014, RS-IMPA-009 and RS-IMPB-006, not only RS-AVR-014. HPLC-04 and impurity-standard certificates are current; method revalidation and RS-AVR-014 evidence are absent.
Note on F6: QC-16 — Correction: same complete chain as Impurity A. RS-IMPB-006 is current to 2028-01-31; MTH-HPLC-012 revalidation and RS-AVR-014 evidence remain absent.
Note on F7: QC-17 — Correction: total impurities uses the same MTH-HPLC-012 chain and all three method-cited standards. Missing method revalidation and RS-AVR-014 evidence prevent use.
Note on F8: QC-18 — OSM-02 calibration and MTH-PHY-009 revalidation are current; no controlled certificate/expiry for required RS-OSM-004 was located.
Note on F9: QC-19 — Correction: MTH-PHY-015 requires RS-PSL-012. PM-01 calibration 2026-12-01, standard expiry 2028-06-02 and method revalidation 2024-08-22 cover planned use.
Note on F10: QC-20 — Method revalidation is current, but no controlled CCI-02 calibration evidence was located. “Current in due month” does not demonstrate qualification on the use date.
Note on F11: QC-21 — MTH-MIC-001 revalidation 2024-01-18 and INC-03 calibration 2026-12-19 are within the appendix intervals; method cites no reference standard.
Note on F12: QC-22 — The appendix sets a 36-month requalification interval. The available 2023 requalification evidence does not cover the March 2027 pull; the 2026 calibration is a separate control.
Note on F13: QC-23 — Correction: GC-01 and KF-01 serve methods outside AVI401-SPEC-v03 and are excluded by the appendix definition; they are not campaign-clearance prerequisites.

## Impurity A Analysis
Lot | Presentation | Data through | Tomas 24 m projection | Tomas 36 m projection | Tomas residual SE | Tomas one-sided upper bound at 36 m | Tomas conclusion | QC 24 m projection | QC 36 m projection | QC residual SE | QC one-sided upper bound at 24 m | QC one-sided upper bound at 36 m | QC supported shelf life / interpretation | QC source locator | QC comment ref
AVI401-R01 | 100 mL | 18 months | 0.31 | 0.46 |  |  | Supports 36 months | 0.37243 | 0.533 | 0.0084092 | 0.39353 | 0.5684 | 24 months supported by current linear model; 36 months not supported because the one-sided 95% upper confidence bound exceeds 0.50%. Confirm with actual 24-month data. | STB-AVI401-002_Long_Term_Results.xlsx, Results (IMPA; months 0–18); AVT-SPEC-018 v03 | QC-25
AVI401-R02 | 250 mL | 18 months | 0.3 | 0.45 |  |  | Worst case; supports 36 months | 0.37071 | 0.525 | 0.0029881 | 0.37821 | 0.53758 | 24 months supported by current linear model; 36 months not supported because the one-sided 95% upper confidence bound exceeds 0.50%. Confirm with actual 24-month data. | STB-AVI401-002_Long_Term_Results.xlsx, Results (IMPA; months 0–18); AVT-SPEC-018 v03 | QC-26
AVI401-R03 | 500 mL | 18 months | 0.36 | 0.49 |  |  | Supports 36 months | 0.44629 | 0.636 | 0.0033806 | 0.45477 | 0.65023 | 24 months supported by current linear model; 36 months not supported because the one-sided 95% upper confidence bound exceeds 0.50%. Confirm with actual 24-month data. | STB-AVI401-002_Long_Term_Results.xlsx, Results (IMPA; months 0–18); AVT-SPEC-018 v03 | QC-27
Note on I2: QC-25 — Ordinary least-squares linear fit to reportable Impurity A values at 0, 3, 6, 9, 12 and 18 months. Formula is retained in-cell.
Note on L2: QC-25 — One-sided 95% confidence bound for the fitted mean: fitted value + t(0.95, df=4) × residual SE × sqrt[1/n + (x−x̄)^2/Sxx]. This is not an individual prediction limit.
Note on N2: QC-25 — Model outputs (approximately): 24 m 0.372, upper bound 0.394; 36 m 0.533, upper bound 0.568.
Note on N3: QC-26 — Controlled source identifies AVI401-R02 as 100 mL, not the retained Tomas baseline 250 mL. Outputs: 24 m 0.371, upper bound 0.378; 36 m 0.525, upper bound 0.538.
Note on N4: QC-27 — Controlled source identifies AVI401-R03 as 250 mL, not the retained Tomas baseline 500 mL. It remains worst case. Outputs: 24 m 0.446, upper bound 0.455; 36 m 0.636, upper bound 0.650.

## Authority Response
Request | Tomas baseline response / evidence | Baseline owner | QC decision | QC basis / correction | Accountable owner | Supporting function | Due / gate | Objective closure evidence | Controlled source locator | QC comment ref
1. P.7 bookmarking and pagination | No QC action | Regulatory CMC | GO | Administrative defect only; QC confirms no new technical content is requested. Regulatory CMC owns corrected P.7 leaf and publishing QC. | Regulatory CMC Manager | Regulatory Operations | 2027-03-19 | Approved correctly bookmarked/paginated P.7 leaf and publishing QC record. | HA-AVI401-IR-2027-014 point 1 | QC-28
2. Impurity A trend and shelf life | Attach refreshed 24-month projections | QC Stability | CONDITIONAL GO | QC can provide an independently checked regression from verified data through 18 months now. Results: all 24-month upper bounds <0.50%; all 36-month upper bounds >0.50%. Present data support 24 months, not 36. Actual 24-month results do not yet exist. R03 is the 250 mL worst case. | QC Stability Scientist | Regulatory CMC | 2027-03-12 for response evidence; refresh after verified 24-month impurities | Approved calculation with source values, formulas, residual SE, one-sided 95% bounds and explicit limitation; later verified 24-month refresh. | Impurity A Analysis I2:P4; STB-AVI401-002_Long_Term_Results.xlsx | QC-29
3. Overwrap and headspace support | Current stability summary confirms control | MSAT / Regulatory CMC | HOLD | Filed leaf supplies BF-51 OTR 2.4 cm³/m²·d·bar at 23 °C/65% RH, but no controlled evidence located for residual headspace oxygen control/measurement in 100 mL and 250 mL fills. Do not infer missing values. | MSAT Engineer | Regulatory CMC | 2027-03-10 | Approved container/pack report with exact page citation for OTR and attributable headspace control/measurement, or explicit confirmed statement that no measurement exists. | AVT-CMC-038 v03 section 5; HA request point 3; ACT-16 | QC-30
4. Accelerated significant change | 0.46% is within specification; no action | QC Stability | HOLD | Significant change and completion of STB-AVI401-003 are supported. The filed leaf names mobile-phase preparation error and confirmatory Impurity B result 0.14%, but the approved LAB-2025-114 investigation and documented assessment of other sequences were not located. | QC Laboratory Investigation Owner | QC Stability; Regulatory CMC | 2027-03-10 | Approved LAB-2025-114 with assignable cause, replacement result and bounded campaign/sequence impact assessment. | AVT-CMC-038 v03 sections 2,4 and data table; HA request point 4; ACT-17 | QC-31
5. Filtration and hold coverage | Validation report references | Validation / MSAT | HOLD | AVT-VAL-104/105 support challenge bounds of 8 h at 30 °C (and 2.1 bar in VAL-104), but no controlled production record or quantitative 47 mm disc versus 254 mm cartridge area-to-mass comparison was located. | MSAT Engineer | Validation Engineer; Regulatory CMC | 2027-03-10 | Approved process record defining maximum production filtration/hold time and temperature, plus Validation-approved conservative area-to-mass calculation with exact report citations. | AVT-VAL-104 pp.3,16; AVT-VAL-105 pp.3,32; HA request point 5; ACT-18 | QC-32
6. Infusion and leachables | Product is already registered | CMC / Regulatory | HOLD | AVT-VAL-203 is approved but the available report does not state the requested infusion duration, administration-set materials, analytical evaluation threshold, or last-fraction leachable result. Registration status is not responsive evidence. | Validation Engineer | MSAT; Regulatory CMC | 2027-03-10 | Approved protocol/report or verified raw-data summary stating duration, set materials, AET and last-fraction result, with page/data locators. | AVT-VAL-203; HA request point 6; ACT-19 | QC-33
7. Invalidated Impurity B run | Replacement result closes the event | QC Stability | DEFER TO COM | The 24-month pull is scheduled for 2027-03-01; no 24-month results exist today. The authority requests verified results as the next update under COM-AVI401-008 and does not request an earlier pull. | Regulatory CMC Manager | QC Stability Scientist | After complete QC release; target 2027-04-29 | Complete verified 24-month dataset including CCI, approved trend assessment and RIM-recorded COM-AVI401-008 handoff/submission decision. | HA-AVI401-IR-2027-014 point 7; Study Status; ACT-20 | QC-34
Note on D2: QC-28 — Administrative defect only; QC confirms no new technical content is requested. Regulatory CMC owns corrected P.7 leaf and publishing QC.
Note on D3: QC-29 — QC can provide an independently checked regression from verified data through 18 months now. Results: all 24-month upper bounds <0.50%; all 36-month upper bounds >0.50%. Present data support 24 months, not 36. Actual 24-month results do not yet exist. R03 is the 250 mL worst case.
Note on D4: QC-30 — Filed leaf supplies BF-51 OTR 2.4 cm³/m²·d·bar at 23 °C/65% RH, but no controlled evidence located for residual headspace oxygen control/measurement in 100 mL and 250 mL fills. Do not infer missing values.
Note on D5: QC-31 — Significant change and completion of STB-AVI401-003 are supported. The filed leaf names mobile-phase preparation error and confirmatory Impurity B result 0.14%, but the approved LAB-2025-114 investigation and documented assessment of other sequences were not located.
Note on D6: QC-32 — AVT-VAL-104/105 support challenge bounds of 8 h at 30 °C (and 2.1 bar in VAL-104), but no controlled production record or quantitative 47 mm disc versus 254 mm cartridge area-to-mass comparison was located.
Note on D7: QC-33 — AVT-VAL-203 is approved but the available report does not state the requested infusion duration, administration-set materials, analytical evaluation threshold, or last-fraction leachable result. Registration status is not responsive evidence.
Note on D8: QC-34 — The 24-month pull is scheduled for 2027-03-01; no 24-month results exist today. The authority requests verified results as the next update under COM-AVI401-008 and does not request an earlier pull.

## Action Ledger
Action ID | Linked sheet / cells | Decision or deficiency | Accountable owner | Supporting function | Due / gate | Objective closure evidence | Controlled source locator | Status | Closure record / next review trigger | QC reviewer
ACT-01 | Operation Gates E2; Qualified Chain F12 | Execute the scheduled pull for all three lots and place samples in controlled quarantine pending ACT-12. | QC Stability Scientist | Stability Operations | 2027-03-01 08:00 | Attributable pull/receipt and quarantine records for R01/R02/R03; identity, condition and time verified; no analytical use. | Study Status, Pull Schedule rows 20:22; MER-2026-4587 | AUTHORIZED / NOT STARTED | QC review when all three receipt/quarantine records are available; testing remains blocked by ACT-12. | Viktor Takacs
ACT-02 | Operation Gates E4; Qualified Chain F8 | Obtain controlled qualification and expiry evidence for RS-OSM-004; osmolality remains held. | QC Reference Standards Coordinator | QC Stability Scientist | Before 2027-03-08 | Effective certificate covering intended use date and documented identity/status check. | MTH-PHY-009 v3.0 section 3 | OPEN — HOLD | QC Stability reviews immediately on certificate receipt; only then may osmolality start. | Viktor Takacs
ACT-03 | Operation Gates E5:E6; Qualified Chain F4:F7 | Obtain controlled qualification and expiry evidence for RS-AVR-014; assay and impurities remain held. | QC Reference Standards Coordinator | QC Method Owner | Before 2027-03-15 | Effective certificate covering all assay/impurity use dates and documented identity/status check. | MTH-HPLC-011/012 v3.0 section 3 | OPEN — HOLD | QC reviews on certificate receipt; assay also needs its pre-use check, impurities additionally require ACT-04. | Viktor Takacs
ACT-04 | Operation Gates E6; Qualified Chain F5:F7 | Obtain approved periodic-revalidation evidence for MTH-HPLC-012. | QC Method Owner | QC Stability Scientist | Before 2027-03-22 | Effective revalidation report showing pass and use interval covering 2027-03-22 to 2027-03-26. | AVT-STB-102-A4 v2.0 section 3; MTH-HPLC-012 v3.0 | OPEN — HOLD | QC reviews on approved record receipt; impurities start only after ACT-03 and ACT-04 close. | Viktor Takacs
ACT-05 | Operation Gates E8; Qualified Chain F10 | Obtain controlled CCI-02 calibration evidence covering planned execution. | QC Instrument Owner | QC Stability Scientist | Before 2027-04-13 | Effective passing calibration certificate covering 2027-04-13 to 2027-04-14. | MTH-CCI-003 v3.0; AVT-STB-102-A4 v2.0 section 3 | OPEN — HOLD | QC reviews on certificate receipt; CCI remains held until acceptance. | Viktor Takacs
ACT-06 | Operation Gates E3 | Execute appearance and pH only after ACT-12 closes and after documented pre-use chain verification. | QC Stability Scientist | QC Analytical | 2027-03-04 | Valid reportable results and independent review for all three lots. | AVT-SPEC-018 v03; applicable methods | BLOCKED BY ACT-12 | QC review when all six results and raw-data review are complete. | Viktor Takacs
ACT-07 | Operation Gates E4 | After ACT-12 closes, execute particulate matter; do not start osmolality until ACT-02 also closes. | QC Stability Scientist | QC Analytical | 2027-03-12 | Three valid particulate results; osmolality results only after documented hold release. | AVT-SPEC-018 v03; MTH-PHY-015 v3.0 | PARTIAL AUTHORIZATION | QC review at particulate completion and separately at ACT-02 closure. | Viktor Takacs
ACT-08 | Operation Gates E5 | Execute assay only after ACT-12 and ACT-03 close. | QC Stability Scientist | QC Analytical | 2027-03-18 | Three valid assay results with system suitability and independent review. | AVT-SPEC-018 v03; MTH-HPLC-011 v3.0 | BLOCKED BY ACT-12/03 | QC review immediately after ACT-03 closure and again when all assay results are reviewed. | Viktor Takacs
ACT-09 | Operation Gates E6 | Execute impurity panel only after ACT-12, ACT-03 and ACT-04 close. | QC Stability Scientist | QC Analytical | 2027-03-26 | Nine valid IMPA/IMPB/IMPT results with system suitability and independent review. | AVT-SPEC-018 v03; MTH-HPLC-012 v3.0 | BLOCKED BY ACT-12/03/04 | QC review after both prerequisites close and after all impurity results are reviewed. | Viktor Takacs
ACT-10 | Operation Gates E7 | Execute sterility only after ACT-12 closes, with required controls and incubation. | QC Microbiology Manager | QC Stability Scientist | 2027-03-16 | Three reportable sterility results, valid controls and independent review. | AVT-SPEC-018 v03; MTH-MIC-001 v3.0 | BLOCKED BY ACT-12 | QC review at completed incubation and second-person review. | Viktor Takacs
ACT-11 | Operation Gates E8 | Execute CCI only after ACT-12 and ACT-05 close. | QC Stability Scientist | QC Instrument Owner | 2027-04-14 | Three valid CCI results with system suitability and independent review. | AVT-SPEC-018 v03; MTH-CCI-003 v3.0 | BLOCKED BY ACT-12/05 | QC review after calibration acceptance and again when all CCI results are reviewed. | Viktor Takacs
ACT-12 | Qualified Chain F12; Operation Gates E2:E8 | Resolve CHM-25-03 qualification interval and impact on samples; common gate for all reportable testing. | Validation Engineer | Stability Operations; QA; QC Stability | Before any reportable testing | Approved current requalification evidence covering the relevant storage period, or approved deviation/impact assessment with sample disposition and documented authorization to test. | AVT-STB-102-A4 v2.0 section 3; CC-2023-004; AVT-VAL-316; MER-2026-4587 | OPEN — ALL TESTING HOLD | QC reviews immediately on approved evidence; only a documented acceptable sample disposition releases test-specific gates. | Viktor Takacs
ACT-13 | Impurity A Analysis I2:P4 | Independently verify the 18-month linear regression and one-sided 95% confidence bounds. | QC Statistics Reviewer | QC Stability Scientist | 2027-03-12 | Signed calculation review reproducing source values, formulas, residual SE and bounds. | Impurity A Analysis I2:M4; Long_Term_Results.xlsx Results | OPEN | QC disposition review on signed calculation verification. | Viktor Takacs
ACT-14 | Operation Gates E9 | Reconcile and approve complete 24-month dataset. | QC Stability Scientist | QC Analytical; QC Microbiology | Target 2027-04-16 | 30/30 required lot-test results verified; investigations resolved; approved completeness record. | AVT-SPEC-018 v03; AVT-SOP-024 sections 7-8 | BLOCKED BY ACT-02:12 | QC review after the last required result, expected after CCI. | Viktor Takacs
ACT-15 | Authority Response D3 | Prepare point 2 response using verified 18-month analysis; explicitly separate projection from unavailable 24-month observations. | QC Stability Scientist | Regulatory CMC | 2027-03-12 | Approved text/table with ACT-13 review and statement that current model supports 24 months, not 36. | HA request point 2; Impurity A Analysis | OPEN | Regulatory handoff review when ACT-13 and draft text are complete. | Viktor Takacs
ACT-16 | Authority Response D4 | Resolve overwrap OTR citation and headspace oxygen evidence gap. | MSAT Engineer | Regulatory CMC | 2027-03-10 | Approved source/page for OTR and headspace control/measurement, or attributable confirmation that measurement does not exist. | HA request point 3; AVT-CMC-038 v03 section 5 | OPEN — RESPONSE HOLD | QC reviews the cited evidence or explicit unavailable-data statement on receipt. | Viktor Takacs
ACT-17 | Authority Response D5 | Provide approved LAB-2025-114 cause, replacement result and analytical-campaign scope assessment. | QC Laboratory Investigation Owner | Regulatory CMC | 2027-03-10 | Approved investigation with assignable cause, 0.14% replacement-result traceability and assessment of other sequences. | HA request point 4; AVT-CMC-038 v03 section 4 | OPEN — RESPONSE HOLD | QC reviews when the approved investigation and scope conclusion are supplied. | Viktor Takacs
ACT-18 | Authority Response D6 | Provide production filtration bounds and Validation-approved scale-down area-to-mass comparison. | MSAT Engineer | Validation Engineer; Regulatory CMC | 2027-03-10 | Approved production record plus quantitative 47 mm disc/254 mm cartridge comparison demonstrating conservativeness. | HA request point 5; AVT-VAL-104/105 | OPEN — RESPONSE HOLD | QC reviews when both process and validation evidence are available. | Viktor Takacs
ACT-19 | Authority Response D7 | Provide administration-set duration/materials, AET and last-fraction leachable result. | Validation Engineer | MSAT; Regulatory CMC | 2027-03-10 | Approved protocol/report or verified data summary with exact locators for all four requested facts. | HA request point 6; AVT-VAL-203 | OPEN — RESPONSE HOLD | QC reviews when complete attributable evidence is supplied. | Viktor Takacs
ACT-20 | Operation Gates E11; Authority Response D8 | Prepare COM-AVI401-008 verified 24-month update after complete dataset release. | Regulatory CMC Manager | QC Stability Scientist | Target 2027-04-29 | ACT-14 closed; approved 24-month tables/trends including CCI; RIM owner records handoff/submission decision. | HA request point 7; Study Status Study B8 | BLOCKED BY ACT-14 | QC hands off only after ACT-14 closure; Regulatory then performs the attributable RIM review. | Viktor Takacs

## QC Disposition
Decision field | QC entry
Disposition | CONDITIONAL APPROVAL — scheduled pull into controlled quarantine may proceed. All reportable testing is held pending storage-chain resolution; test-specific holds also apply. Dataset and unsupported response claims remain held.
Decision date | 2027-02-08
Reviewer | Viktor Takacs
Physical pull decision | GO only to pull AVI401-R01/R02/R03 on 2027-03-01 and place them in controlled quarantine. Pull does not release samples for testing; ACT-12 controls that release.
Testing decisions | ALL TESTING initially HOLD pending ACT-12 (CHM-25-03 qualification interval/sample impact). After ACT-12: conditional GO for appearance, pH, particulate matter and sterility; osmolality also needs ACT-02, assay ACT-03, impurities ACT-03/04, and CCI ACT-05.
Complete dataset release | HOLD. Release only after all specification-required 24-month results (30 total), including CCI, are reportable, independently reviewed and reconciled; target QC review 2027-04-16.
Written-response evidence handoff | CONDITIONAL GO by 2027-03-19 for verified evidence then available. Exclude ungenerated 24-month results; close or explicitly state gaps for headspace oxygen, LAB-2025-114 scope, filtration scale-down/process bounds and administration-set details.
COM-AVI401-008 handoff | HOLD until complete 24-month QC release. Target handoff 2027-04-29; Regulatory CMC owns the RIM decision/update.
Open action IDs by gate | Common pull/test gate: ACT-01 and ACT-12. Test-specific gates: ACT-02 to ACT-11. Dataset gate: ACT-14. March response: ACT-13 and ACT-15 to ACT-19. COM-AVI401-008: ACT-20.
Evidence basis | HA-AVI401-IR-2027-014; STB-AVI401-002 Study Status and Long-Term Results; AVT-SPEC-018 v03; AVT-STB-102-A4 v2.0; cited method/qualification/standard records; AVT-VAL-104/105/203; AVT-SOP-024 and AVT-SOP-072.
Next attributable QC review trigger | First attributable review is on Validation/QA-approved ACT-12 evidence before any reportable testing. Subsequent reviews occur at each test-specific closure, on 2027-03-19 for the response pack, and after the last verified result for dataset/commitment release.
Decision statement | QC authorizes the scheduled pull into quarantine. No reportable test may begin until the chamber qualification/sample-impact gate is closed; thereafter only tests whose full chain is evidenced may proceed. No due-month assumption, incomplete dataset, or unavailable measurement may be represented as acceptable evidence.
Note on B2: QC-35 — Final QC operating position. Baseline recommendations remain visible in their original cells for traceability but are not adopted where contradicted by this disposition.